## Purchase list
Echo Yeast hardware module — final purchase list
Updated 2026-10-14 · Items already ordered and newly ordered · Unit prices are the amounts actually paid
No. | Category | Item | Specification / model | Qty | Unit price (CNY) | Subtotal (CNY) | Purpose
1 | Structural parts & printing consumables | 3D printing filament (PLA) | PLA Basic, white 1000 g + black 1000 g (without spool) | 1 spool each | 77.38 | 77.38 | FDM printing of the collection cup, the heater-block body and the three plug-in modules
2 | Temperature control | Heat-resistant thermal silicone pad | 200 × 400 × 0.3 mm, 6 W/(m·K) | 1 sheet | 25 | 25 | Thermal interface between the heating film and the heater block; cut to size
3 | Temperature control | PI heating film | 60 × 70 mm, 12 V / 25 W | 2 pcs | 24.64 | 24.64 | Bonded to the underside of the heater block as the heating element
4 | Temperature control | MAX31865 module + PT100 | MAX31865 converter module + PT100 platinum RTD (3-wire) | 1 set | 33.5 | 33.5 | Temperature acquisition: PT100 resistance converted to a digital temperature by a 16-bit ADC
5 | Temperature control | KSD9700 thermal cut-off | 60 °C, 250 V / 10 A, metal case, normally closed | 1 pc | 6.12 | 6.12 | Hardware-level over-temperature protection; in series with the heater so it physically cuts the circuit
6 | Temperature control | MOSFET driver module | IRF520 | 1 pc | 3.29 | 3.29 | Takes the PWM signal and drives the heating film to implement PID power control
7 | Temperature control | Cooling fan | 40 × 40 × 10 mm, 12 V, quiet | 1 pc | 10 | 10 | Mounted on the heat sink to keep the electronics underneath and the block from overheating
8 | Temperature control | Heat sink | Aluminium, 70 × 36 × 100 mm | 1 pc | 23 | 23 | Dissipates heat downwards and protects the circuit underneath
9 | Optics | LED light source | 485 nm high-brightness LED | 3 pcs | 17.34 | 17.34 | Excitation light source; spares kept for wavelength margin
10 | Optics | LED constant-current driver | Adjustable constant-current driver board, 350 / 500 mA | 1 pc | 36.38 | 36.38 | Provides a stable current for the 485 nm LED and reduces intensity drift
11 | Optics | Photodiode | BPW34, through-hole package | 4 pcs | 8.1 | 8.1 | Detects the 528 nm fluorescence at the receiver end; spares kept
12 | Optics | Narrow-band filters | Ø10 mm, 30 nm bandwidth, OD3; 485 nm × 2 + 528 nm × 2 | 4 pcs | 154 | 154 | One at each end separates the excitation light from the fluorescence spectrum
13 | Optics | Op-amp | OPA350UA (SOIC-8) | 2 pcs | 18.8 | 18.8 | Forms the transimpedance amplifier with the photodiode and outputs a measurable voltage
14 | Optics | Light shield (black heat-shrink tubing) | 12 mm bore, 10× thick wall | 1 pack | 8 | 8 | Wraps the emitter and receiver plugs to exclude ambient light
15 | Optics | Focusing lenses for emitter and receiver | Matched to the 485 nm emitter and the 528 nm receiver (newly ordered) | 2 sets | TBD | TBD | Focus the LED output onto the centre of the vessel and the fluorescence onto the photodiode active area
16 | Electronics & signal | Microcontroller | Arduino Nano V4 (board + USB cable) | 1 pc | 62.84 | 62.84 | Timing control, ADC readout and the PID temperature algorithm
17 | Electronics & signal | ADS1115 high-precision ADC | ADS1115 module, 16-bit, 4-channel differential, I²C (newly ordered) | 1 pc | TBD | TBD | Reads the amplified photodiode voltage with high precision
18 | Electronics & signal | SOT89 adapter board | SOT89 to DIP through-hole adapter (newly ordered) | TBD | TBD | TBD | Lets SOT89 SMD parts be soldered through-hole
19 | Electronics & signal | SOT223 adapter board | SOT223 to DIP through-hole adapter (newly ordered) | TBD | TBD | TBD | Lets SOT223 SMD parts be soldered through-hole
20 | Power & wiring | Power adapter | 24 V / 2 A (48 W), with mains lead | 1 pc | 12.87 | 12.87 | Supplies the instrument; 24 V bus
21 | Power & wiring | LM2596 buck module | LM2596S DC-DC adjustable step-down | 1 pc | 4.49 | 4.49 | Steps 24 V down to 12 V (heating film, fan) and 5 V (logic)
22 | Power & wiring | Breadboard / protoboard | Single-sided, rosin-coated green, 12 × 18 cm | 2 pcs | 10.87 | 10.87 | Prototyping during debugging, then soldered permanently
23 | Power & wiring | Jumper wires | 21 cm, male–male / male–female / female–female, 120 pcs total | 1 set | 9.87 | 9.87 | Temporary connections between boards and modules
24 | Power & wiring | Electrolytic capacitors | 16 V / 100 µF, 5 × 11 mm | 2 pcs | 6.74 | 6.74 | Supply filtering; suppresses switching noise
25 | Power & wiring | Resistor kit | 1/4 W metal film, 25 common values, 500 pcs total | 1 set | 7.9 | 7.9 | Current limiting, voltage division and transimpedance feedback
26 | Power & wiring | Indicator LEDs | 5 mm red / green / yellow, 60 pcs total | 1 set | 3.46 | 3.46 | Power and run-status indication
27 | Power & wiring | Terminal blocks | KF301-2P × 10 + 3P × 10, 5.08 mm | 1 pack each | 7.17 | 7.17 | Power input and load wiring
28 | Power & wiring | 30 AWG silicone wire | Extra-flexible silicone wire for wiring inside the plugs and for signals (newly ordered) | TBD | TBD | TBD | Fine wiring inside the plugs keeps them compact and flexible
29 | Tools & consumables | Thermal silicone adhesive | Kafuter K-5204K, fast-curing, white | 1 tube | 13.2 | 13.2 | Fixes the heating film, PT100 and heat sink; also seals against light
30 | Tools & consumables | Heat-resistant double-sided tape | PI polyimide tape, 0.1 mm thick, 3 mm × 10 m | 1 roll | 10.7 | 10.7 | Holds the heating film onto the heater-block surface
31 | Tools & consumables | Silicone wire | 22 AWG red + black, rated ≥ 105 °C | 1 roll each | 13.05 | 13.05 | Power wiring for the heater and supply circuits
32 | Tools & consumables | Desoldering wick | For desoldering and rework (newly ordered) | 1 roll | TBD | TBD | Rework and clearing of mis-soldered joints
33 | Tools & consumables | Multimeter | UNI-T UT33A, auto-ranging | 1 pc | 73 | 73 | Voltage, current, continuity and resistance measurement
34 | Tools & consumables | Soldering station (936) | 60 W adjustable constant temperature, standard model | 1 set | 66 | 66 | Soldering the circuits and plug-in modules
35 | Tools & consumables | Hot-melt glue gun | 40 W, with glue sticks | 1 pc | 15.35 | 15.35 | Securing components and sealing light leaks
36 | Tools & consumables | Solder wire | Lead-free, 63% tin, 0.8 mm, rosin core | 1 roll | 27.65 | 27.65 | Soldering consumable
|  | Total, items already ordered (CNY) |  |  |  | 790.71 | Newly ordered items will be added once prices are known

## Summary by category
Summary by category (items already ordered)
Category | Items | Amount (CNY) | Share
Structural parts & printing consumables | 1 | 77.38 | 9.79%
Temperature control | 7 | 125.55 | 15.88%
Optics | 7 | 242.62 | 30.68%
Electronics & signal | 4 | 62.84 | 7.95%
Power & wiring | 9 | 63.37 | 8.01%
Tools & consumables | 8 | 218.95 | 27.69%
Total | 36 | 790.71 | 100%

## Purchase links
Purchase links (items already ordered)
Item | Link
3D printing filament (PLA) | https://e.tb.cn/h.8D2q9Wgd9FICQlA?tk=4wliTOpQjq4
Heat-resistant thermal silicone pad | https://e.tb.cn/h.8D2u35e03HzrnKQ?tk=ZUlpTOpnJNj
PI heating film | https://e.tb.cn/h.8DdYQSZ9I7Gn80H?tk=q5L3TOpvEUn
MAX31865 module + PT100 | https://e.tb.cn/h.8Eo6m9Y1yBcqsCB?tk=0hdiTOpzeHk
KSD9700 thermal cut-off | https://e.tb.cn/h.8EoisstnhcDy4WH?tk=nJM9TOKZxsu
MOSFET driver module | https://e.tb.cn/h.8FTxa7A424Jc4TE?tk=Gf2cTOKc66T
Cooling fan | https://e.tb.cn/h.8FTCMdiqBypUWOx?tk=15evTOKX21L
Heat sink | https://e.tb.cn/h.8vh4852tGrpl6rI?tk=MU7NTOKf3Em
LED light source | https://e.tb.cn/h.8vhUy8oFTunUWog?tk=3NKaTOKg7rV
LED constant-current driver | https://e.tb.cn/h.8EoqXI9Ckjpxpg9?tk=B6LBTOK6QOH
Photodiode | https://e.tb.cn/h.8vhSus9evlimDIM?tk=drwcTOK7u8T
Narrow-band filters | https://e.tb.cn/h.8F6ngurIWsopc7N?tk=susnTOKFY6Y
Op-amp | https://e.tb.cn/h.8DWYYiqtEKvqeLx?tk=O7y0TOKxbDm
Light shield (black heat-shrink tubing) | https://e.tb.cn/h.8ELgpJKO9UjsSkU?tk=7ORATOKzjE0
Microcontroller | https://e.tb.cn/h.8ELBloVll81AYYP?tk=ZfVFTOqJpe2
Power adapter | https://e.tb.cn/h.8F6tQHNABFojNUa?tk=Vds9TOKARZj
LM2596 buck module | https://e.tb.cn/h.8DWdMqYnWoZeTiV?tk=qyQATOqZxlt
Breadboard / protoboard | https://e.tb.cn/h.8DWeeX1yDMySMdW?tk=gWEGTOqbyTD
Jumper wires | https://e.tb.cn/h.8vS0UJczgCS6B9I?tk=UVU1TOqX7Bc
Electrolytic capacitors | https://e.tb.cn/h.8DWTRQb5qqQUXFB?tk=XnXFTOq3l20
Resistor kit | https://e.tb.cn/h.8Fh0TDbu7fD8gpo?tk=QrEITOq4cHV
Indicator LEDs | https://e.tb.cn/h.8ELnMVRucfBtMhc?tk=ielnTOqUEXm
Terminal blocks | https://e.tb.cn/h.8ELo5JbhvB3KreD?tk=uO7KTOqgam2
Thermal silicone adhesive | https://e.tb.cn/h.8Fhd5G7Prf3J7S0?tk=t5SDTOqT7Sk
Heat-resistant double-sided tape | https://e.tb.cn/h.8DWQ2ns5GLIGnhN?tk=KnPSTOqSk57
Silicone wire | https://e.tb.cn/h.8ELJRdCyEZp1FDH?tk=5qyfTOq7oBz
Multimeter | https://e.tb.cn/h.8Fh5C8Nsg2KiLLM?tk=q5DbTOqjBmh
Soldering station (936) | https://e.tb.cn/h.8DWpiXAZNefdTRU?tk=A0vmTOqObkV
Hot-melt glue gun | https://e.tb.cn/h.8ELFziG8HafPdXK?tk=Mr0HTOqNcMG
Solder wire | https://e.tb.cn/h.8FhLCVc3yUmjPZa?tk=Lg8xTOquJTR

## To be completed
Items still to be completed in this list
No. | Item to be completed | Notes / what we need from you
1 | Prices of the newly ordered items | The amounts actually paid and quantities for the six new items — SOT89 and SOT223 adapter boards, ADS1115 module, emitter/receiver focusing lenses, 30 AWG silicone wire and desoldering wick. The total updates once these are filled in.
2 | Purchase links for the newly ordered items | If the list should keep the purchase links as well, send me the Taobao links for these six items and I will add them to the "Purchase links" sheet.
3 | Specifications of the focusing lenses | Focal length, aperture and material for the emitter and receiver lenses (for example a φ6 mm / f8 mm plano-convex lens), so the key parameters can be fixed in the technical requirements and the lenses reordered later.
4 | Actual model of the LED constant-current driver board | The link in the list is a generic adjustable board. If a specific model or measured output current range has been settled on, please add it.
5 | Specific part numbers of the SOT89 / SOT223 devices | The adapter boards themselves have no model requirement, but if the SMD parts to be soldered (a regulator or a MOSFET, for instance) are already known, note them down so the package compatibility can be confirmed.
6 | Number and wiring of the PI heating films | The list has 2 pcs at 12 V / 25 W. Please confirm whether they are used in parallel or bonded to the two sides of the block, so the technical requirements can describe them consistently.
7 | Vendor and quotation for the CNC parts | A vendor has been contacted to machine the test parts. If a quotation or machining list is available, it is worth a separate "outsourced structural machining" sheet, counted apart from the component purchases.
8 | Whether the purchase categories should follow "structure / temperature / optics / electronics / power / tools" | The list currently uses these six categories. If the university’s reimbursement rules require a different breakdown, tell me and I will adjust it.