Technical Requirements — Temperature-controlled Optical Readout Module

Echo Yeast Hardware Module (current revision)

Revision: 14 October 2026　　Prepared by: Echo Yeast Hardware Team

1　Scope

This specification defines the terms, technical requirements, test methods, inspection rules, and marking, packaging, transport and storage requirements for the hardware module of the Echo Yeast whole-cell biosensor.

The module holds and incubates a liquid sample containing the yeast sensor (cuvette or our in-house collection cup), provides a 485 nm excitation path and a 528 nm fluorescence collection path, and outputs raw measurement data for recording. It is intended for laboratory validation and small-batch prototypes.

2　Normative References

GB/T 38517-2020　Particles — General rules for bioaerosol sampling and analysis

GB/T 1804-m　General tolerances — Tolerances for linear and angular dimensions without individual tolerance indications

GB 9706.1　Medical electrical equipment — Part 1: General requirements for basic safety

ISO 10993　Biological evaluation of medical devices

WS 648-2019　General hygienic requirements for air disinfection machines (applicable when used as an add-on module)

3　Terms and Definitions

Terms defined in GB/T 38517-2020 apply. In this document, "module" refers to the unit formed by the heated block, the detachable optical plugs, the electronics and the power supply.

4　Technical Requirements

4.1　Overall Performance

| Item | Requirement |
| --- | --- |
| Purpose | Temperature-controlled incubation and fluorescence readout of the enriched liquid sample from airborne influenza A virus aerosols; provides a stable temperature environment and an optical readout path for the yeast whole-cell biosensor |
| Sample vessel | Test stage: standard 10 mm path-length cuvette. Project stage: in-house 3D-printed collection cup |
| Working volume | Cuvette type ≥ 1.5 mL; collection-cup type 20 mL liquid medium (with yeast) |
| Set point | 30 °C (yeast culture temperature) |
| Temperature accuracy | ±0.5 °C (steady state, within 10 min) |
| Temperature uniformity | Top-to-bottom difference of the liquid in the vessel ≤ 0.5 °C |
| Heat-up time | Room temperature (25 °C) to 30 °C in ≤ 5 min |
| Excitation wavelength | 485 ± 5 nm |
| Emission wavelength | 528 ± 10 nm |
| Optical channel | Single channel; excitation and collection axes at 90° |
| Target | HA protein on the surface of influenza A virus (H1N1), driving yEGFP expression through the GPCR pathway |
| Detection type | Qualitative (present / absent); signal expressed as relative fluorescence intensity |
| Readout time | ≤ 10 s per reading (including settling delay); incubation and induction follow the experimental protocol |
| Power supply | DC 24 V / 2 A adapter; stepped down internally to 12 V and 5 V |
| Power consumption | ≤ 30 W peak during heating; ≤ 10 W steady state |
| Module dimensions | Heated block plus optical plugs ≤ 120 × 90 × 80 mm (excluding external power supply and cabling) |
| Weight | ≤ 1.5 kg for the module itself (excluding the adapter) |

4.2　Structural Requirements

4.2.1　Heated Block and Vessel Interface

| Item | Requirement |
| --- | --- |
| Print material | PLA (white / black); inner surfaces painted matte black or lined with light-absorbing fabric to suppress stray light |
| Print process | FDM, layer height 0.1–0.2 mm, wall thickness ≥ 2 mm, infill ≥ 20% |
| Dimensional accuracy | Mating surfaces ±0.2 mm; general outline ±0.5 mm |
| Heated block body | Rectangular block with two Ø vessel bores (cuvette / collection cup), bore depth ≥ 30 mm; optical channels machined on the side faces |
| Optical channel | 8 mm bore; channel axis perpendicular to and intersecting the vessel bore axis; coaxiality of the two ends ≤ 0.3 mm |
| Vessel fit | Clearance to the vessel wall ≤ 0.3 mm; filled with thermal pad or a thin sleeve where necessary |
| Detachable plugs | Emitter_Plug, Receiver_Plug and Heater_Plug are independent plug-in parts; no play once inserted; removal must not affect the block |
| Plug locating | Cylindrical surface plus shoulder; no play after 20 insert / remove cycles |
| Light sealing | Optical channel and plug-to-block joints must be fully light-tight; plugs wrapped in ≥ 12 mm black heat-shrink tubing, gaps sealed with black silicone or hot glue |
| Light-tightness | With the excitation source off under room lighting, receiver output drift ≤ 1% of full scale |
| Heat dissipation | Aluminium heat sink (70 × 36 × 100 mm) and 40 × 40 × 10 mm, 12 V quiet fan mounted under the block |
| Assembly | Plugs and vessel can be inserted and removed without tools; M3 self-tapping holes or heat-set inserts |
| General tolerance | Per GB/T 1804-m; unspecified fillets R0.5, unspecified chamfers C0.5 |

4.2.2　Detachable Optical Plugs

Emitter_Plug, Receiver_Plug and Heater_Plug are independent plug-in parts. Emitter_Plug carries the 485 nm LED, the focusing lens and the excitation filter; Receiver_Plug carries the emission filter, the focusing lens and the BPW34 photodiode; Heater_Plug routes the heating-film leads and the PT100 into the block. The plug-to-block interface is identical across revisions, so switching vessels (cuvette ↔ in-house collection cup) requires no redesign of the plugs.

4.3　Temperature Control

| Item | Requirement |
| --- | --- |
| Heating element | PI polyimide heating film, 60 × 70 mm, 12 V / 25 W, two pieces |
| Mounting | Bonded with PI-based high-temperature thermal tape (0.1 mm); edges fixed and sealed with K-5204K thermal adhesive |
| Sensor | PT100 platinum resistor (Class A, 3-wire), inserted into a blind bore filled with thermal grease |
| Transmitter | MAX31865 module, SPI, 16-bit ADC, resolution ≤ 0.03 °C |
| Control range | Room temperature to 40 °C |
| Set point | 30 °C |
| Accuracy | ±0.5 °C (steady state) |
| Power drive | IRF520 MOSFET driver module, PWM with PID algorithm |
| Over-temperature (hardware) | KSD9700 thermal fuse (60 °C, 10 A, metal case, normally closed) in series with the heater for physical cut-off |
| Over-temperature (software) | Software threshold ≤ 45 °C; PWM output disabled immediately when exceeded |
| Cooling | 40 mm, 12 V quiet fan plus aluminium heat sink |
| Thermal interface | 6 W/(m·K) thermal pad (200 × 400 × 0.3 mm, cut to size) and K-5204K thermal adhesive |
| Insulation | Insulation resistance between heater leads and metal body ≥ 2 MΩ |

4.4　Optical Detection

| Item | Requirement |
| --- | --- |
| Excitation source | 485 nm high-brightness LED with adjustable constant-current driver (350 / 500 mA) |
| Excitation focusing | Focusing lens in the emitter end converging the LED output to the centre of the vessel; spot diameter ≤ 3 mm |
| Excitation filter | Ø10 mm narrow-band filter, CWL 485 nm, FWHM 30 nm, OD3 |
| Emission filter | Ø10 mm narrow-band filter, CWL 528 nm, FWHM 30 nm, OD3 |
| Emission focusing | Focusing lens in the receiver end converging fluorescence onto the photodiode active area |
| Photodetector | BPW34 silicon PIN photodiode (through-hole) |
| Geometry | Excitation and collection axes at 90°, both perpendicular to and intersecting the vessel bore axis at the same height |
| Transimpedance stage | OPA350 single-channel op-amp as transimpedance amplifier; feedback resistor selected to match signal level, gain adjustable |
| Light sealing | Emitter and receiver ends wrapped in black heat-shrink tubing; joints to the block sealed against light |
| Alignment | After insertion, the intersection of the optical axes deviates from the vessel bore axis by ≤ 0.5 mm |
| Serviceability | LED, filters, focusing lenses and photodiode are detachable and can be replaced or cleaned individually |
| Detection limit | Against a blank medium without HA as negative control, target signal-to-noise ratio S/N ≥ 3; the actual limit is determined experimentally |

4.5　Signal Acquisition and Temperature-control Unit

| Item | Requirement |
| --- | --- |
| Core board | Arduino Nano development board; handles sequencing, ADC readout and the temperature-control algorithm |
| ADC | ADS1115 high-precision ADC, 16-bit, 4 differential channels, I²C, programmable gain |
| Temperature readout | PT100 resistance read via SPI from the MAX31865 and converted to temperature |
| Power output | PWM output to the IRF520 driver module to set heater power |
| Acquisition | Switches the 485 nm LED on and off on a fixed timing sequence, reads the receiver voltage and logs background and signal continuously |
| Data output | Timestamp, temperature and raw ADC counts emitted over USB / serial in a fixed format for offline analysis and archiving |
| Resolution | Effective number of bits ≥ 14; background noise ≤ 2 LSB |
| Sampling rate | Temperature ≥ 1 Hz; optical signal ≥ 10 Hz (adjustable) |
| Time base | On-board crystal oscillator; drift ≤ 1 s over a continuous run of ≥ 1 h |
| Power distribution | 24 V input stepped down by LM2596 to 12 V (heater, fan) and 5 V (logic) |

4.6　Power Supply and Auxiliaries

| Item | Requirement |
| --- | --- |
| Input | DC 24 V / 2 A adapter (48 W) with 5.5 × 2.1 mm barrel connector |
| Step-down | LM2596S adjustable DC-DC module, 24 V to 12 V and 5 V |
| Filtering | 100 µF / 16 V electrolytic capacitors placed close to each supply node |
| Circuit carrier | 12 × 18 cm single-sided perfboard for soldered assembly; breadboard for the debugging stage |
| Power wiring | 22 AWG silicone wire (red / black), rated ≥ 105 °C, for heater and supply loops |
| Signal wiring | 30 AWG extra-flexible silicone wire for wiring inside the plugs and between boards |
| Board-to-board | 21 cm Dupont jumpers (male-male, male-female, female-female) |
| SMD adapters | SOT89 and SOT223 adapter boards for through-hole assembly of SMD parts |
| Passives | 1/4 W metal-film resistor kit (25 common values, 500 pieces); 5 mm red / green / yellow indicator LEDs |
| Terminals | KF301-2P / 3P, 5.08 mm pitch, interlocking |
| Soldering tools | 936 constant-temperature soldering station (60 W), 0.8 mm rosin-cored lead-free solder, desoldering wick |
| Measurement | UNI-T UT33A auto-ranging digital multimeter |
| Fixing and sealing | 40 W hot glue gun and glue sticks for securing parts and sealing light paths |
| Routing | Power and signal wiring routed separately to keep switching noise away from the microampere-level photocurrent signal |

4.7　Safety and Biosafety

| Item | Requirement |
| --- | --- |
| Electrical safety | In line with the relevant clauses of GB 9706.1: metal parts grounded, heater and signal circuits routed separately |
| Over-temperature | KSD9700 hardware fuse and software threshold as dual protection |
| Over-current | Fuse or resettable fuse on the supply side; current-limiting element in the heater loop |
| Surface temperature | Touchable surface ≤ 45 °C; insulating shroud over the heated face |
| Biosafety | Any work with live virus must be carried out in a laboratory of the appropriate biosafety level with suitable PPE, per GB/T 38517-2020 clause 4.5.4 |
| Sample handling | Collected liquid samples are stored and transported per GB/T 38517-2020 clause 4.6.7 (0–10 °C, delivered within 10 h) |
| Disinfection | Sampling head, vessel and operators' hands disinfected before and after sampling, per GB/T 38517-2020 clause 4.5.5 |
| Waste | Waste liquid and disposables containing biological samples are handled as infectious medical waste and never discharged to drains |
| Sterilisation compatibility | Vessel and plugs tolerate 75% ethanol wiping; autoclaving requires prior verification that PLA parts do not deform |

4.8　Firmware Requirements

PID temperature control: PWM duty adjusted from the PT100 reading to hold the vessel bore at 30 ± 0.5 °C;

Optical acquisition: the 485 nm LED is switched on according to a fixed timing sequence, the photodiode voltage is read after a settling delay and converted into a relative fluorescence intensity;

Sequence logging: timestamp, temperature and raw ADC counts written out in a fixed format for offline analysis and archiving;

Protection logic: heating output is disabled and an error flag set immediately on over-temperature, sensor open circuit or short circuit.

5　Test Methods

| Item | Method |
| --- | --- |
| Temperature accuracy | Set 30 °C; after 20 min stabilisation, measure the vessel bore with a calibrated thermometer (or a second PT100); deviation ≤ ±0.5 °C |
| Temperature uniformity | Measure at several heights inside the vessel bore; difference ≤ 0.5 °C |
| Heat-up time | Record the time from room temperature to 30 °C; ≤ 5 min |
| Over-temperature protection | Break the temperature feedback so heating runs continuously; confirm the KSD9700 cuts the heater circuit before 60 °C |
| Optical alignment | Insert and remove the plugs 10 times; check spot position repeatability with a paper screen or visible light; deviation ≤ 0.5 mm |
| Light-tightness | With the excitation source off under room lighting, acquire for 5 min; output drift ≤ 1% of full scale |
| Background and noise | Fill with blank medium without yeast, acquire continuously; compute mean and standard deviation of the background |
| Fluorescence response | Load HA-induced and uninduced yeast samples; compare readings on the 528 nm channel |
| Repeatability | Measure the same positive sample ≥ 5 times; relative standard deviation RSD ≤ 5% (target, to be confirmed by data) |
| Re-assembly consistency | Disassemble and reassemble the vessel and plugs 5 times; reading difference ≤ 5% |

6　Inspection Rules

First-article inspection: dimensional report, vessel-to-block fit record and optical alignment record for the first unit of each batch;

Routine inspection: temperature accuracy, light-tight background, optical alignment and insulation resistance for every module;

Type inspection: full inspection whenever a new design is finalised or the structure, optical layout or heating scheme changes.

7　Marking, Packaging, Transport and Storage

Marking: model, name, supply rating, manufacturer and date on the module; function and insertion direction marked on each plug;

Packaging: protective film on optical surfaces, plugs packed separately, dust-, scratch- and static-proof; inspection record enclosed;

Transport: avoid high temperature, humidity and strong vibration;

Storage: 0–40 °C, relative humidity ≤ 80%, no corrosive gases.

8　Notes

This module is the readout unit paired with the aerosol sampling front end, and can also be used on its own for laboratory validation of the yeast sensor.

This revision covers the three core functions of temperature control, optics and signal acquisition; the user interface and host software lie outside its scope and will be specified separately in a later stage.

All unspecified tolerances per GB/T 1804-m; unspecified fillets R0.5, unspecified chamfers C0.5.

The optical channel bore, filter centre wavelengths and focusing-lens focal lengths are the key parameters of this revision; any change requires the light-tightness and alignment tests to be repeated.

The "working volume" and "response time" clauses reflect the latest experimental data: working volume is ≥ 1.5 mL for the cuvette type and 20 mL liquid medium for the collection-cup type; under ideal laboratory conditions a 528 nm fluorescence signal is detectable 5–15 min after the sample is added, and the response time of the complete module is to be verified experimentally.